SIL O2

Configuration Settings for Factory Pre-programming

END CUSTOMER:________________________

| Alarm | Alarm value (% Oxygen) | Hysteresis (% Measuring range) | Delay time (s) | Alarm type | Function |
| --- | --- | --- | --- | --- | --- |
| Alarm 1 (relay-1) |  |  |  | Choose an item. | Choose an item. |
| Alarm 2 (relay-2) |  |  |  | Choose an item. | Choose an item. |
| *Alarm 3 (Digital) |  |  |  | Choose an item. | Choose an item. |

Please complete the above table according to your requirements the instructions below.

An alarm delay up to a maximum of 9.9 seconds guarantees that briefly exceeded alarm values do not activate the alarm.

Available alarm types:

MAX alarm 	with rising measured value

MIN alarm 	with falling measured value

Available alarm functions:

De-energized:	in good status the relay is powered off

Energized:	in good status the relay is powered on

No function: 	alarm output is switched off permanently

Hysteresis:   Difference in the points at which the relay operates when traversing an alarm level      setpoint.  i.e  if  a normally energised relay De-energises on a rising alarm at the setpoint, it can be set to Re-energise at a value below that setpoint when falling.(units in % of measuring range. Thus for the standard measuring range of 0-25%, a 1.0% Hysteresis value as shown be, would result in a % volume Oxygen Hysteresis value of 0.25% Vol. Oxygen Hysteresis value. ((25%x1.0%)/100.=0.25%). Example below.

Hysteresis in % of range (0-25%)

Relay 1 Min

Alarm @ 19.5%

Healthy@20.5%

In the case on both these relays, Energised is the Healthy condition (closed circuit) whereas the Alarm condition causes an Open circuit.

4%

Relay 2 Max

Alarm @ 5%

Healthy@4.75%

1%

% Oxygen

0%        5%            10%             15%          20%       25% 25%

*Alarm 3 is a Transistor output. This is not to be confused with Relay 3 (RL3) which is the main safety Trip Relay for the SIL O2 Module and which cannot be configured.

Customer:

Date:

Rev.3